Word Wizard

Educator Reference Manual

Setup • Daily Operation • Care • Safety

Ohio Northern University – Team 5  |  2026

# Introduction

Word Wizard is a portable, NFC-based augmentative and alternative communication (AAC) device designed to support non-verbal individuals in educational settings. When a communication card is tapped on the device, Word Wizard speaks the corresponding word or phrase aloud using a built-in text-to-speech (TTS) engine.

This manual covers everything a teacher or classroom staff member needs to set up, operate, maintain, and safely care for the device.

# Safety Warnings

⚠️ SCREWS / FASTENERS The enclosure uses accessible screws to allow component servicing. Do not allow unsupervised individuals to open the device. Loose screws are a choking and injury hazard.

⚠️ BATTERY SAFETY Do not expose the device to extreme heat, puncture the enclosure, or submerge in liquid. Use any USB-C cable and a standard 5V charger. If the device becomes hot to the touch, smells unusual, or shows signs of swelling, discontinue use immediately.

⚠️ DROP RISK Although the TPU bumper provides impact resistance for drops up to approximately 3 feet onto hard flooring, repeated or severe impacts can still damage internal components. Inspect the device after any significant drop.

# What’s in the Box

- Word Wizard device (with TPU bumper attached)

- Communication card set (color-coded per grammatical group)

- Card storage drawer

📌 NOTE Additional communication cards can be programmed using a smartphone with NFC write capability. See Section 5 for details.

# Setup & First Use

## Charging the Device

- Connect the USB-C cable to the port on the back of the device.

- Plug the other end into a standard USB wall adapter (5V recommended).

- If powered on, green and orange status indicator lights will confirm charging is in progress. If the lights are blinking, disconnect charger and discontinue use.

- Allow a full charge before first use. The battery capacity is at least 40 Watt-hours, providing extended daily use on a single charge.

📌 TIP Charge overnight before the first school day. We recommend charging each evening to ensure a full battery every morning.

## Powering On

- Toggle the power switch on the side.

- The device will boot and be ready to use in less than a minute.

- Place the device on a flat, stable surface accessible to the student.

# Daily Operation

## Using Communication Cards

Each communication card contains an embedded NFC tag. When a card is tapped on the top surface of the device, Word Wizard will speak the corresponding word or phrase aloud.

- Hold the card flat and touch it to the NFC reader area, indicated by the magic wand logo.

- The device will respond within a moment and speak the word.

- Return the card to the storage drawer when not in use.

## Card Color Coding (based on the Fitzgerald Key)

Cards follow the Fitzgerald Key color coding system, a widely recognized AAC standard that reduces cognitive load by grouping words by grammatical category:

| Color | Category | Examples |
| --- | --- | --- |
| Yellow | Question Words | why, where, how |
| Orange | Action Words | hungry, help, stop |
| Green | Social Words | yes, no, thank you |
| Pink | People | Mom, brother, grandma |

## Volume

The device is calibrated to output 60–65 dB at ear level, consistent with normal speech.

## Language Support

Word Wizard uses the Piper TTS engine, which supports multiple languages and dialects. Contact the ONU Engineering Department if language or dialect settings need to be changed for a specific student.

# Programming & Customizing Cards

One of the key advantages of Word Wizard is that communication cards can be customized or reprogrammed without purchasing new hardware. You will need a smartphone with NFC write capability (most modern Android and iOS devices support this).

- Download a free NFC writing app on your smartphone (e.g., “NFC Tools” for iOS or Android).

- Open the app and select “Write”, then “Add a Record”, then “Text”.

- Enter the word or phrase you want the card to speak.

- Select “Write / X Bytes” and hold the blank or existing NFC card against the back of your phone until the app confirms a successful write.

- Tap the newly programmed card on Word Wizard to verify it works correctly.

📌 NOTE NFC programming requires a smartphone capable of NFC write functionality. The device must be powered on to play a card, but does not need to be connected to a network. Cards can be reprogrammed as many times as needed.

# Charging & Battery Care

- Use only standard USB-C cables rated for 5V charging.

- Avoid charging in direct sunlight or in environments above 35°C (95°F).

- Do not leave the device fully discharged for extended periods.

- A rechargeable internal battery means no disposable batteries are needed, reducing waste.

- If the device will not be used for more than one week, store it with a partial charge (around 50%).

⚠️ DO NOT Do not use non-standard or high-wattage fast chargers not rated for the device. Do not attempt to replace the battery yourself; contact the ONU Engineering Dept.

# Cleaning & Hygiene

Word Wizard is designed with school hygiene in mind. The PLA enclosure and TPU bumper can be wiped down regularly without degradation.

## Recommended Cleaning Method

- Power off the device before cleaning.

- Use a cloth lightly dampened with water and/or a gentle cleaning solution.

- Avoid using alcohol to clean the acrylic back plate to reduce the risk of solvent stress cracking.

- Wipe all exterior surfaces, including the bumper and NFC reader area.

- Allow to dry fully before powering back on.

- The communication cards can be wiped with the same method.

⚠️ AVOID Do not submerge the device in any liquid. Do not spray cleaners directly onto the device. Do not use abrasive materials or bleach-based cleaners, which may damage surface materials.

📌 NOTE Surface pores in 3D-printed parts may accumulate debris over time. Consistent wiping with cleaning solutions helps minimize bacterial buildup.

# Communication Card Storage

The device includes an integrated card storage drawer at the base, organized by a modified Fitzgerald Key color coding. 40+ cards can be stored in the drawer.

- Keep cards sorted by color category for easy retrieval.

- Return cards to the drawer after each use to minimize loss.

- Cards should be stored flat to prevent bending or damage to the NFC chip.

- Consider labeling the drawer sections by color group if students manage cards independently.

📌 TIP Use the pictogram or image printed on the card face to help students identify and select cards independently, reducing reliance on reading ability.

# Troubleshooting

| Problem | Likely Cause | Solution |
| --- | --- | --- |
| Device does not power on | Battery depleted | Charge device for at least 30 minutes before retrying. |
| Card tap produces no sound | Card not aligned / NFC issue | Re-tap the card flat and centered on the reader area. Try a different card to isolate the issue. |
| Sound is too quiet | Grill obstruction | Ensure speaker grill is clear of debris. Wipe grill with a dry cloth. |
| Card speaks the wrong word | Card was reprogrammed incorrectly | Re-write the card using the NFC programming steps in Section 5. |
| Device feels hot during charging | Charging environment too warm | Move device to a cooler location. If heat persists after moving, discontinue charging. |
| Device does not charge | Cable or port issue | Try a different USB-C cable or charger. Inspect port for debris and clear gently with a dry toothpick if needed. |

📌 STILL HAVING ISSUES? If the problem is not resolved by the steps above, do not attempt to open the device enclosure. Contact your building administrator or the team responsible for device maintenance.

# Quick Reference Card

Cut out or print this section and post it near the device for easy access.

| ✅ Daily Startup Checklist □ Device charged overnight □ Power on–wait for “Powered On” vocal indicator □ Cards sorted in drawer by color □ Device placed within student’s reach □ Test one card tap before class begins | ⚠️ Key Reminders • USB-C only for charging • Wipe with damp cloth, never submerge • Do not open the enclosure • Report issues to ONU Engineering Dept. |
| --- | --- |